## others_en
|  | Observatory La Silla, Chile | PROJEKT: PS 1257 PROJECT: PS 1257 | REVIZE: REVISION : | 0.
|  | Servisní návštěva a kontrolní protokol Service visit and check protocol | DOKUMENT:1257_03_06.xx DOCUMENT:1257_03_06.xx | DATUM : DATE : | 05/2013
Checked: 2016 November 6-7, P. Pravec, K. Hornoch, M. Velen
point | describe |  |  | status - comment |  | frequency
A - oil system
1. | visual check oil level in tank (basement into pier) |  |  | OK |  | 6 months
2. | hearing of running oil pumps (smooth working) |  |  | OK (minor leakage, considered insignificant) |  | 6 months
3. | check tightness of oil system (tank, pumps) |  |  | OK |  | 6 months
4. | check tightness of oil system into pier (on observing floor) pipes, pressure probe |  |  | OK |  | 6 months
5. | check tightness of oil system - telescope covers on DEC,RA |  |  | OK |  | 6 months
B - dome
1. | check of state measure tapes around dome (cleanness, adhesion) |  |  | OK (some minor grime cleaned) |  | 6 months
2. | check of state moving contacts on brush wear, pollution (if needed lubricated) |  |  | not OK, needs to be fixed |  | 6 months
3. | check of state ventilation windows (visual and hearing) on smooth working |  |  | OK |  | 6 months
4. | check of rubber collar around dome (if needed lubricated) |  |  | OK, except for a hole on the north side, to be sealed |  | 6 months
5. | check of state solar panels mechanical, on pollution (if needed clean it by soap water) |  |  | OK (some chickweed washed) |  | 12 months
6. | check of clamping pyrgeometer PG1, in case of pollution cleaneed by soap water |  |  | OK (minor filth cleaned) |  | 12 months
7. | check of outside cable trace on mechanical damage and UV degradation |  |  | OK |  | 12 months
8. | visual and hearing check of dome rotating during 360° working (smooth working) |  |  | OK |  | 12 months
9. | check of GPS NTP and GSM antennas outside (mechanical, pollution), if needed clean it by soap water |  |  | the GSM antenna remains broken (since 2014) |  | 12 months
C - slit
1. | check of slit proximity switches - down S27,S28,S47 (distance swith - reflect surface 2-4mm) |  |  | S47 adjusted, S27 and S28 work too (S27 for closed position) |  | 6 months
2. | check of slit proximity switches - up outside S29, S30 (distance swith - reflect surface 2-4mm) |  |  | both switches adjusted |  | 6 months
3. | visual check inside of R5 slit cabinet and bateries on integrity |  |  | OK |  | 6 months
4. | visual and hearing check of slit moving during opening and closing (smooth working) |  |  | OK |  | 6 months
5. | check of state bateries (dome not on contact, sky without sun, value on display GU1 not less than 21,6V |  |  |  |  | 6 months
D - telescope
1. | check of flexibil pipe with cables about telescope montage on integrity |  |  | OK |  | 6 months
2. | visual and hearing check telescope engines (flaps,DEC, HA,autoguider,shutter, filters) smooth moving |  |  | OK (autoguider was not checked) |  | 6 months
E - meteo station
1. | visual and hearing check rotating of anemometer (smooth running) |  |  | OK |  | 6 months
2. | check allsky camera lens on pollution if needed clean it by soap water |  |  |  |  | 6 months
3. | check state of overvoltage protection in MC1 cabinen on mast (BR0 signalling must not light red) |  |  | OK |  | 6 months
4. | check of integrity protection pipe from meteostation to telescope building (can be demage from animals) |  |  | OK |  | 12 months
F - others
1. | check of UPS (display without failure text, fans are working smoothly) |  |  | OK |  | 18 months
2. | check state of overvoltage protection in R1 (BR1 signalling must be green, BR2,BR3 must not light red) |  |  | OK |  | 6 months
3. | check state of overvoltage protection in RH1 behind control room door (BR-MC1 signalling must green) |  |  | OK |  | 6 months
4. | check state connection cables to hand paddles on integrity |  |  | OK |  | 12 months